## Slide 1: WBPM Test Results and Paper Work Status
2025. 10. 17.
송우진

## Slide 2: Introduction
Experimental Objective
  To test the performance of the WBPM developed for measuring the white-beam profile at the 4GSR beamline.
  To evaluate the WBPM performance in measuring the beam profile under different gas environments and with varying beam positions and sizes.
Experimental Details
  Measurement of beam profiles under different gas environments.
    Residual gas, Xenon gas (@ 1 X 1 mm pink beam)
  Performance test according to beam size changes
    1.2 x 1.0 mm -> 1.0 x 0.5 mm -> 0.9 x 0.2 mm (@ center)
  Performance test according to beam displacement
    7 steps, beam size 1 x 1 mm

## Slide 3: Schematic Diagram
Multi-Pole Wigger
Ec = 12 keV
Normal source size
: 500 um x 30um
Total power : 10kW
36m
Double Multilayer
Monochromator
Ru/C, 4.0 nm
: 10 – 20 keV
W/B4C, 2.6 nm
: 20 – 55 keV
DCM 1
Si (220)
: 30 – 50 keV
DCM 2
Si (111)
: 12 – 33 keV
Sample stage
Air bearing rotation stage
sCMOS Camera
Pco.Edge 5.5
Hamamatsu Fusion
2x ,4x, 5x, 10x, 20x, 40x
Slit
LMM
Diamond
1.4 mm
Al window
0.1 mm
IVU24
2.4 x 2.4 mrad (H x V)
Sigma (H) : 12.9 urad
Power=17.94 kW
DCM
WBPM
(0m)
Shielding
wall
(21m)
(28m)
(14m)
Fixed Mask
300 x 300 urad
Power= 8.25 kW
Movable Mask
77 x 77 urad
(1.6 x 1.6 mm)
Power= 699 W
White beam Slit
33 x 33 urad
(1 x 1 mm)
Power= 208 W
(24.5m)
(30m)
IVU24
White beam Slit
Fixed Mask
Movable Mask
Application scenario at Korea-4GSR
PLS-II BL6C test - Pink beam (Total reflection beam by DMM 1st crystal)

## Slide 4: 1st Beam Image
Verification of availability in residual gas environments (Hard X-ray)
Residual gas (middle -7 torr)
Xenon injection (middle -6 torr)
Exposure time: 0.3 sec

## Slide 5: Image Pre-processing
Screen: Use the projection result without additional processing
WBPM: Background image removal and image rotation
Absolute color-bar scale
Exposure time: 75 sec
X-ray burn paper

## Slide 6: Calculation of pix/mm using the active area
Center: (397.0, 293.5) pixel
Diameter: 692.356 pixels = 24 mm
pix./mm = 34.664 um/pix
Threshold=100000

## Slide 7: Beam Size Measurement

## Slide 8: Beam Size Measurement
Beam Size: 1.2 x 1.0 mm, 1.0 x 0.5 mm, 0.9 x 0.2 mm
Horizontal
Vertical
1.2 x 1.0 mm
1.0 x 0.5 mm
0.9 x 0.2 mm

## Slide 9: Beam Position Measurement
Position data was analyzed shot by shot  0.3 sec exposure time, total of 600 shots.
The position resolution: Deviation of the 600 image centers & Fitting error  29.37 µm.
The center position was obtained using Gaussian fitting.

## Slide 10: Beam Position Measurement

## Slide 11: WBPM Test Result
The WBPM test was carried out at PLS-II BL6C.
The PSF is predicted to be 120 μm or less during beam profile measurement
The position measurement resolution is 29 μm based on an exposure time of 0.3 ms.

## Slide 12: Paper Preparation
Dissertation: Entire content written
- Excluded: Beam size measurement experiment
- Included: Magnification defocusing field design, experiment, and correction method
IBIC proceeding: Beam size measurement experiment
- Excluded: Design, Result analysis, Magnification
- Included : Photo of the whole equipment, WBPM measurement images
Journal: Resolution calculation and verification (Experimentally)
- Excluded: Magnification, Design detail
- Included: Resolution, result analysis
Internal report (Planned): Technical design, operation method
- Excluded: Theory (Resolution, Principle detail, …)
- Included: Fabrication process, PLS-II experiment details

## Slide 13: Paper Preparation
Introduction
Principle and configuration
Monitor resolution
Experimental set-up
Measurement
Conclusions
Logical Structure
Necessity:
DLSR  ε↓  Requires high position measurement accuracy.  Conventional devices cannot withstand power density.  Non-destructive measurement via IPM method.
Central cone at 10 urad level used in undulator beamline.  Requires profile measurement to identify the central cone.  Requires precise electric field design (resolution).
Features:
1. Resolution determining factor (PSF) analysis via particle tracking simulation.
2. Verified at PLS-II beamline.

## Slide 14: Paper Preparation
Introduction
Principle and configuration
Monitor resolution
Experimental set-up
Measurement
Conclusions
+
+
+
+
+
-
-
-
-
-
-
Extraction field
Repeller electrode
Guiding electrode
Photo-electrons
Photo-ions
Readout
+
Beam
Microchannel plate (MCP)
Phosphor screen
Camera
Visible light

## Slide 15: Paper Preparation
Introduction
Principle and configuration
Monitor resolution
Experimental set-up
Measurement
Conclusions

## Slide 16: Paper Preparation
Introduction
Principle and configuration
Monitor resolution
Experimental set-up
Measurement
Conclusions
Factors | Sources | Quantity () |  | 
 |  | PAL-XFEL SX | Korea-4GSR IVU | PLS-II BL6C
Field quality | Non-linearity | Under 4 μm |  | 
 | Defocusing field | (Magnification) |  | None
Transverse spread | Thermal energy | Kr, 70 – 120 μm (Dep. on beam E) | Xe, 55 μm (IVU20) | Xe, 51 μm (MPW)
 | Momentum transfer | Under 0.1% |  | 
 | Space charge effect | None |  | 
Readout device | MCP resolution | 80 (~ 100) μm |  | 
Data processing | Calibration method | None (single axis) | 10 μm (±3 mm) | None ()
Point Spread Function = |  | 100 – 145 μm | 97.68 μm | 94.96 μm
Factors | Sources |  |  | 
 |  | PAL-XFEL SX | Korea-4GSR IVU | PLS-II BL6C
Field quality | Non-linearity | Under 4 μm |  | 
 | Defocusing field |  |  | None
Transverse spread | Thermal energy | Kr, 70 – 120 μm (Dep. on beam E) | Xe, 55 μm (IVU20) | Xe, 51 μm (MPW)
 | Momentum transfer | Under 0.1% |  | 
 | Space charge effect | None |  | 
Readout device | MCP resolution | 80 (~ 100) μm |  | 
Data processing | Calibration method | None (single axis) | 10 μm (±3 mm) | 
 |  | 100 – 145 μm | 97.68 μm | 94.96 μm

## Slide 17: Paper Preparation
Introduction
Principle and configuration
Monitor resolution
Experimental set-up
  Schematic diagram
  PLS-II BL6C, Xenon gas injection
Measurement
Conclusions
36m
Slit
LMM
Diamond
1.4 mm
Al window
0.1 mm

## Slide 18: Paper Preparation
Introduction
Principle and configuration
Monitor resolution
Experimental set-up
  Schematic diagram
  PLS-II BL6C, Xenon gas injection
Photon beam | Beamline | Event number (Xe) | Event number (Residual gas)
White beam | 4GSR IVU20 | 7.91e+9 | 1.76e+8
 | PLS-II IVU20 | 3.00e+9 | 9.25e+7
 | PLS-II MPW14 | 2.44e+7 | 7.94e+4
10 keV beam (0.1% bW) | 4GSR IVU20 | 2.50e+6 | 8.19e+3
 | PLS-II IVU20 | 1.85e+5 | 6.07e+2
 | PLS-II MPW14 | 7.76e+4 | 2.55e+2
30 keV beam (0.1% bW) | 4GSR IVU20 | 2.59e+4 | 5.24e+1
 | PLS-II IVU20 | 5.09e+2 | 1.03e+0
 | PLS-II MPW14 | 3.75e+3 | 7.60e+0

## Slide 19: Paper Preparation
Introduction
Principle and configuration
Monitor resolution
Experimental set-up
Measurement
Conclusions

## Slide 20: Summary
Analysis of PLS-II BL6C experimental results completed and results sharing meeting concluded.
After completing the logical structure and outline of the paper, currently drafting the content.
Still considering whether to include the PSF measurement experiment conducted at XFEL.