## Slide 1: Validation of an Absolute Coordinate Calibration Device
23. January 2026
Juhwan Yoon

## Slide 2: Index
Overview
Device configuration
Ground truth
Construction method
Boundary condition
Measured image
Image acquisition
Perspective transformation
Analysis
1D pixel sum profile
Difference examination
Tilted screen distorting beam profile
Raw image from camera [px]
Desired transverse image [mm]

## Slide 3: Overview of an experimental setup
가설: 진단카메라 이미지를 빔단면으로 변환하면 빔 프로파일도 함께 바뀜
방법: 빔단면에 해당하는 ground truth를 준비해놓고, 변환된 이미지와 비교
-> 진단 카메라의 스크린 이미지와 빔단면에 투영된 스크린 이미지에서 각각 같은 boundary condition 좌표쌍 지정
-> 빔단면 좌표쌍은 2축 스테이지를 통하여 측정하고, 이 값이 이미지 변환의 정확성 결정
-> 이미지 품질 차이 고려하여 정규화, 스케일 통일 등을 거친 후 1D pixel sum 프로파일 차이값 및 개형 비교
고려사항: ground truth와 진단카메라의 이미지 품질 차이
Camera Space [px]
Real Space [mm]

## Slide 4: Ground truth
Requirements
Beam cross-section is invariable regardless of the screen alignment
Camera and beam are parallel installed

## Slide 5: Boundary condition coordinates
입력 명령 좌표 : (-21.884 mm, 5.879 mm)
입력 명령 좌표 : (-21.572 mm, -24.061 mm)
입력 명령 좌표 : (11.411 mm, 7.819 mm)
입력 명령 좌표 : (11.656 mm, -22.244 mm)

## Slide 6: Coordinate calibration
Front grid plate
Rear grid plate
535 mm
1410 mm
834 mm
입력 명령 좌표 1: (-21.884 mm, 5.879 mm)
입력 명령 좌표 2: (-21.572 mm, -24.061 mm)
입력 명령 좌표 3: (11.656 mm, -22.244 mm)
입력 명령 좌표 4: (11.411 mm, 7.819 mm)
CMOS-Scintillator 거리: 834 mm
실제 위치 좌표 1: (-21.925 mm, 5.918 mm)
실제 위치 좌표 2: (-21.350 mm, -24.044 mm)
실제 위치 좌표 3: (11.869 mm, -22.258 mm)
실제 위치 좌표 4: (11.363 mm, 7.829 mm)
ΔX by horizontal scanning
- Δx: motor error
- Δy: tilted grid plate
ΔX by vertical scanning
- Δx: actuator torsion
- Δy: motor error

## Slide 7: Transverse plane

## Slide 8: Diagnostic camera image

## Slide 9: Perspective-transformed image

## Slide 10: Beam profile analysis: Ideal vs. Transformed
빔단면 투영 이미지 (Ground truth)
원근변환 후 이미지 (검증 대상)

## Slide 11
Beam profile analysis: Ideal vs. Raw
진단 카메라 이미지 (검증 대상)
빔단면 투영 이미지 (Ground truth)

## Slide 12: THANK YOU

## Slide 13: Appendix: Ground truth