## Slide 1: Predicting preventable Interruptions to Operations at NSLS-II In collaboration with SLAC, ANL
Mar.07.2024
Reid Smith : Speaker
Feng Bai : Postdoc researcher
IFCA Beam Dynamics Workshop - MLAPA

## Slide 2
  NSLS-II Accelerator Overview
  Context
    Reliability, Downtime, Types of Faults
    Magnet cooling & water systems
  Development
    Philosophy
    Structure
    Methods
  Application & Use
    Prediction parameters
    Use by System experts, Operations
    Next steps
Contents

## Slide 3
NSLS-II Accelerator Overview

## Slide 4
NSLS-II Accelerator Overview
STORAGE RING PARAMETERS
Ring circumference 792 m (.5 mile)
Ring Energy 3 Gev
Ring Current 500 mA
# Cells 30
Cell design type Double-Bend Achromat
Vertical emittance .008 nm-rad
Horizontal emittance .55 nm-rad
Time between bunches 2 ns
Revolution period 2.64 us
RF frequency 500 mHz
#RF Buckets 1320
# Bunches 1056
Active beamlines 31
Complex scattering 5
Hard X-ray scattering & spectroscopy 6
Soft X-Ray scattering & spectroscopy 12
Structural biology 5
Imaging & Microscopy 6

## Slide 5
Chart: NSLS-II yearly Reliability (%)
| FY17 | FY18 | FY19 | FY20 | FY21 | FY22 | FY23
Dumps | 96.8 | 96.5 | 97.1 | 97 | 96.7 | 95.99 | 95.8
Accelerator Reliability goal defined by DOE as 90%
NSLS-II Internal goal of 95%
NSLS2 regularly runs between 95-97% reliability
  …but 5% of 4800 hrs is still 240 hrs of downtime
Even in a reliable machine, downtime is a tangible burden on users.
Accelerator Reliability
Where can we reasonably gain in reliability?:
Fast Faults -
  Power dips, power supply trips, network outages, vacuum spikes, RF trips, cryo quenches, etc.
  Not especially predictable – actionable by system improvement, or reducing trip-recovery
Slow Faults
  Temperature, pressure, flow rate, ground current, etc.
  Trip usually initiated by Equipment Protection systems
  Actionable when caught early enough
How early?
One 2-day maintenance & one 1-day interlocks period w weekend studies, repeating every 3 weeks.
10 days minimum, 20 days preferred
Reliability is % uptime when in Scheduled Ops.

## Slide 6
Issue Summary: local clogs in magnet cooling-water system sporadically cause heating issues with magnets. Seen most in QM magnets.
Normal temps ~35C; Klixons will dump beam between 70-80C unless overridden; Experts will abort at 80C.
Overheating can cause permanent damage to magnets – leading long & costly replacement.
Not the first time we’ve dealt with this issue
Magnet Temps & Cooling water system
# of magnet cooling-water flushes, 3-month bins
Thermal camera image of a hot magnet.

## Slide 7
Plot of quadrupole magnet temperatures in several ‘bad’ spots around the ring; rising over 10 days in January 2024.
Summary-alarm values of Magnet RTD temperatures. Alarms sound in the accelerator Control Room, email experts, etc.
Monitors:
  Plots (visual inspection)
  Alarms
  Calculating change over time (drift)
  Programmatic prediction…
Drift monitor shows difference between reference temp & current temp – colors indicate size of delta
1
2
3
Magnet Temps: the run-up to Predictive Programs

## Slide 8: Machine Learning Operations
Development Philosophy
The name MLOps incorporates machine learning (ML), software development (DEV) and operations (OPS).
Machine Learning Operation is designed to incorporate machine learning engineering and DevOps to seek for data governance and deployment in production with reliability and scalability.
Slide content provided by Feng Bai

## Slide 9
Data Structure for QM magnets
Data for all QM magnets around the ring handled in 4 layers:
Temperature data as a list of HDF5 files
30 Cells in NSLS-II storage ring.
4 sensors/magnet * 4-5 QM/cell = 20-24 PVs per cell
Temp data time series (288 meas/day @ 5 minute sampling rate) for each sensor
Slide content provided by Feng Bai

## Slide 10: Data Engineering and Storage Systems
Data Engineering includes:
1.) Data acquisition based on the network paths from online sensors.
2.) Raw data written into HDF5 files.
3.) Data transfer by reorganizing the raw data into self-structured data in .py files (we reorganize the sensor data into 5 minutes per point in time series)
Data Structure for QM magnets
Q: What to do about asynchronous triggered data? A: data sampled from archiver at fixed interval (5 min by default)
Slide content provided by Feng Bai

## Slide 11: Machine Learning Systems (ML systems)
The input temperature data is retrieved from the archiver appliance, organized, processed, and the outputs are written to files and PVs.
There are three components in the ML systems that are parallelized:
Data Cleaning
Model/Learning
Prediction
Data Structure for QM magnets
Slide content provided by Feng Bai

## Slide 12: 1. Data Cleaning
Outliers
Temperature
Time
Sensor Resolution
Temperature
Time
𝜀
Outlier Detection (Linear regression case)
Sensor resolution (Linear regression case)
regression model without outliers.
regression model including outliers.
regression model accounting for sensor resolution.
regression model without considering sensor resolution.
Data Structure
Filtering out outliers from sensor faults using isolation forest algorithm (over Euclid norms)
Adjusting sensitivity to account for jitter “created by” sensor resolution (0.125 C for our 1-wire sensors)
Slide content provided by Feng Bai

## Slide 13: 2. Modeling & Regression
We use time-series data per PV make temperature prediction in the pipelines. We tried several regression methods and deep learning methods for online predictions.
1.) Linear Regression;
2.) Quadratic Regression;
3.) Nonlinear Exponential Regression;
4.) LSTM.
Methods
3. Prediction
Predictions/Parameters
ALL models:
Days to 40 C
T 3 days
T 7 days
T 14 days
T 30 days
Linear Fit:
Slope
Least sq error
Quadratic Fit:
Square coeff
Linear coeff
Least sq error
LSTM:
<shape parameters>
<error parameter(s)>
The newly trained ML models are used to model future behavior and predict temperatures at points in future.
Slide content provided by Feng Bai

## Slide 14
(Tested but Unused Prediction Models)
Piecewise Linear Model
Overlap Piecewise Linear Model
Linear Models, Lasso/ Ridge
Dataset of temperature for every day:
D: daily dataset;
N: # of daily data inputs (5 mins);
n: number of days;
The regression done every day with N data inputs;
The piecewise linear models can capture the daily parameter (slopes)
Dataset of temperature for every day:
D: daily dataset;
N: # of daily data inputs (5 mins);
n: number of days;
The regression done every 12hs with N data inputs (including the dash lines);
The piecewise linear models can capture the daily parameter (slopes)
1. Lasso
2. Ridge
+λ1
+λ
Strong Regularization
L1 Norm
Smooth Regularization
L2 Norm
Strong Regularization can feature the small parameters with the L1 norm distance to 0. However, Smooth Regularization just feature the characteristics of parameters gradually.
X: data inputs (daily time series temps), : output parameter
Slide content provided by Feng Bai

## Slide 15: Machine Learning Methods: LSTM
The data structures in the LSTM are shown in distributed pipelines per PV, per cell.
The data in the training are displayed in tensor networks with the given period to form the “time window”.
(Tested but Unused Prediction Models)
Slide content provided by Feng Bai

## Slide 16
Results of Fitting
Slide content provided by Feng Bai
Shown: results of Linear/least-squares and Lasso models over a 10-day span.
In production: training and prediction periods are independent.
Better predictions with longer fitting dataset, EXCEPT:
Not able to deal with abnormalities in training set

## Slide 17
Using linear fit prediction right now.
Still testing learning-model fitting and prediction.
Predictions/Parameters
Both models:
Days to 40 C
T in 3 days
T in 7 days
T in 14 days
T in 30 days
Linear Fit:
Slope
Least sq error
LSTM:
<shape parameters>
<error parameter(s)>
Predictions & Displays

## Slide 18
5 C
5 C
Use by Experts
Presently, EE & Mechanical groups are watching hot magnets.
Slope, temp predictions, and ‘time to threshold’ PVs are directing and replacing the same calculations being done by hand or in Excel, with much less time committed.

## Slide 19: Strengths and limitations
+
Slope is a VERY helpful early warning.
Linear prediction catches magnets warming from ‘normal baseline’.
Changing time windows & thresholds can help adapt to situations.
Python; able to add features
-
Fitting models deal poorly with fast changes (repairs, clogs);
Even with perfect modeling, ‘sudden temp jumps’ in modeling time-window will result in bad predictions
LSTM learning model needs more data (both length & type) to be worth using
Next Steps
Expand ML model to use years of history, & other magnet/water readbacks
Decide if all QM temp datasets can be treated (& trained) interchangeably
Consider multivariable Transformer vs LSTM
Can we make a model to deal with sudden temp rise/drop separately?
Applied to QM Magnets
Need to expand scope to make learning model worthwhile:
  If ML is wanted, train using GPU cluster (for parallel processing)
Make the display easier for experts to use/scan:
Future Systems
Power supply ground currents
Possible use in RF and/or Cryo
Next steps & Future applications
Conclusions & Next Steps

## Slide 20: Thanks to: BNL: Feng Bai, Guimei Wang, Yoshi Hidaka, Jinhyuk Choi, Wing Louie, Joseph Gormley, Brian Walsh, Michael Charumaneeroj SLAC: Daniel Ratner, Finn O’Shea, Anwesha Das ANL: Louis Emery, Nikita Kuklev, Ihar Lobach
Mar.07.2024